## test result
TECHMAN ROBOT INC. | [26 empty] | Product Document
Subject: | [6 empty] | MPRM DVT Test Report | [20 empty] | Rev:
Project Code: | [6 empty] | TMAB | [13 empty] | Doc. No.: |  |  |  | MR-405 |  |  | 1
Model Name: | [6 empty] | TM16S | [13 empty] | Release Date: |  |  |  | 2022-04-20 |  |  | 3 Pages
Collision test of human body parts
Purpose:
To confirm DUT(Device Under Test) impact force meets ISO/TS 15066 specification.
Environment:
1. Test Tools: GTE
2. Operation Mode: Collaborative Mode.
3.Safety setting : default
4. Payload: X Y axis 0.67kg、Z axis 0.12kg
Criteria:
1. The DUT should meet the specification.
2. For the maximum force specifications for various parts of the human body, reference to ISO / TS 15066.
3. Static force must be less than half of the maximum force
Condition:
1. Arm Model name & SN:
2. ControlBox Model name、SN& HW Ver:
Medium & heavy payload series standard control box XXXX#CCXXXXXX
3. HMI Ver:
4. Power Manager Ver:
5. Safety Monitor Ver:
6. Joint FW Ver:
7. I/O Ver:
8. Test Date:
9. Tester:
Conclusion
Illustration of this test :
GTE
‧The spring constant of GTE K2 complies with the specifications of ISO 15066
‧The K1 damping elements is are used to simulate the biomechanical properties of body surfaces.
and protect K2 from damage caused by collision.
Illustration of this test :
1.The distance between the target point and the robot base is half reach.
2.Impact tool length is 100 mm.
3.The impact compression value is 50mm.
4.Tool origin point:
| Angle( °) |  |  |  |  | J1 |  |  | J2 |  |  |  | J3 | J4 |  | J5 |  | J6
| CA2206007 | [7 empty] | -50.6 |  |  |  | -135.7 | 97.3 |  | -90
5.Collision velocity setting:
| K2 |  |  | Part |  |  |  | K1 |  | Force Spec. |  | Test Velocity(mm/s) |  |  | Human-machine speed limit(mm/s)
| 10 |  |  | Abdomen |  |  |  | SH10 |  | 220 |  | 162 |  |  | 225
| 25 |  |  | Chest |  |  |  | SH70 |  | 280 |  | 135 |  |  | 195
| 30 |  |  | UpperArm |  |  |  | SH30 |  | 300 |  | 243 |  |  | 315
| 35 |  |  | Back |  |  |  | SH30 |  | 420 |  | 162 |  |  | 225
| 40 |  |  | LowerArm |  |  |  | SH70 |  | 320 |  | 243 |  |  | 315
| 50 |  |  | Thigh |  |  |  | SH30 |  | 440 |  | 162 |  |  | 225
| 60 |  |  | LowerLeg |  |  |  | SH30 |  | 260 |  | 72 |  |  | 125
| 75 |  |  | Hands |  |  |  | SH70 |  | 280 |  | 90 |  |  | 145
↑For HW5.0
Test Result:
Robot S/N |  |  |  |  | #CA2206007 | [18 empty] | Unit: Newton (N)
[11 empty] | X_Axis | [5 empty] | Y_Axis |  |  | Z_Axis | [5 empty] | Result
K2 |  |  | Part |  |  |  | K1 |  | Force Spec. |  | Force (Max) |  |  | Force (static) |  |  | Force (Max) |  | Force (static) | Force (Max) |  | Force (static) |  |  |  | Pass/ Fail
10 |  |  | Abdomen |  |  |  | SH10 |  | 220 | [19 empty] | 0
25 |  |  | Chest |  |  |  | SH70 |  | 280 | [19 empty] | 0
30 |  |  | UpperArm |  |  |  | SH30 |  | 300 | [19 empty] | 0
35 |  |  | Back |  |  |  | SH30 |  | 420 | [19 empty] | 0
40 |  |  | LowerArm |  |  |  | SH70 |  | 320 | [19 empty] | 0
50 |  |  | Thigh |  |  |  | SH30 |  | 440 | [19 empty] | 0
60 |  |  | LowerLeg |  |  |  | SH30 |  | 260 | [19 empty] | 0
75 |  |  | Hands |  |  |  | SH70 |  | 280 | [19 empty] | 0
CONFIDENTIAL
Observation # | [6 empty] | Observation Description | [15 empty] | Issue Reference #
1
2
3
4
K1
(SH10)
K2
K1
(SH70)
K2
K1
(SH30)
K2

## Force photo
| X-Asix | Y-Asix | Z-Asix
10 Abdomen
25 Chest
30 Upper Arm
35 Back
40 Lower Arm
50 Thigh
60 Lower Leg
75 Hands